Spectrometer SDK Documentation

A detailed and comprehensive control setting document for zolix spectrometer device

- spec_get_dll_ver

  - description:

      - get library version number.

  - declaration:

      - void	spec_get_dll_ver(char * buff, int len);

  - arguments:

      - buff : data buffer, 16 bytes at least, before the buffer can be used, it needs to be initialized.

      - len : buffer size.

  - return:

      - void.

- spec_set_usb_mode

  - description:

      - Set the wiring mode between the spectrometer and the host, default is USB mode.

  - declaration:

      - void	spec_set_usb_mode(bool usb);

  - arguments:

      - usb :

        - True - Square-port USB cable connection

        - False - RS232 to USB cable connection

  - return:

      - void.

- spec_enum_dev_count

  - description:

      - get numbers of devices. Available in USB wire connection only.

  - declaration:

      - int	spec_enum_dev_count();

  - arguments:

      - no argument.

  - return:

      - number of device.

- spec_enum_dev_sn

  - description:

      - get information for specified device. Available in USB wire connection only

  - declaration:

      - bool	 spec_enum_dev_sn(int index, char * buff, int len);

  - arguments:

      - index : device index number.

      - buff : data buffer for device information, 16 bytes at least, before the buffer can be used, it needs to be initialized.

      - len : buff size.

  - return:

        - true: get information successfully.

        - false: cannot get information.

- spec_open

  - description:

      - Connect the device according to the device serial number or serial port number, and return the device ID after successful connection.

  - declaration:

      - int	spec_open(const char * sn);

  - arguments:

      - sn : buffer for device serial number,16 bytes at least, before the buffer can be used, it needs to be initialized.

      - len : buff size.

  - return:

      - return device ID number. Invalid if number is less than zero.

- spec_close

  - description:

      - close device according to ID.

  - declaration:

      - bool	spec_close(int id);

  - arguments:

      - id : ID number to be closed.

  - return:

        - true: close device successfully.

        - false:close failed.

- spec_get_is_open

  - description:

      - check if the device is connected.

  - declaration:

      - bool	spec_get_is_open(int id);

  - arguments:

      - id : ID number to be checked.

  - return:

        - true: device is already on.

        - false: the device is off or ID number is not correct.

- spec_get_error

  - description:

      - get device error nformation according to ID number.

  - declaration:

      - bool  spec_get_error(int id, char * buff, int len);

  - arguments:

      - id : device ID number.

      - buff : data buffer for device error information.16 bytes at least, before the buffer can be used, it needs to be initialized.

      - len : buff size.

  - return:

        - true: get error information successfully.

        - false: the device has no error or ID number is not correct.

- spec_set_timeout

  - description:

      - Set the timeout period of the device's read and write operations. Default is 20s.

  - declaration:

      - bool	 spec_set_timeout(int id, int ms_read, int ms_write);

  - arguments:

      - id : ID number.

      - ms_read: set the timeout period of device's read operation.

      - ms_write:set the timeout period of device's write operation.

  - return:

        - true: set complete.

        - false: set fail.

- spec_set_dev_info

  - description:

      - set device system info, including manufacturer, model, serial number, date and firmware version.

  - declaration:

      - bool	spec_set_dev_info (int id, spec_info info);

  - arguments:

      - id : device ID number.

      - info: data structure for sys information. Its properties:

        - manufacturer

        - model

        - sn

        - date

        - firmware_ver

  - return:

        - true: set sys info successfully.

        - false: set fail.

- spec_get_dev_info

  - description:

      - get device system info, including manufacturer, model, serial number, date and firmware version.

  - declaration:

      - bool	 spec_get_dev_info (int id, spec_info * ptr);

  - arguments:

      - id : device ID number.

      - ptr: pointer to data structure for sys info.

  - return:

        - true: get sys info successfully.

        - false: get sys info fail.

- spec_backup

  - description:

      - backup all parameters of the device, including Turret, Grating, Zero positon, Correcting coefficients,  Grating lines and Grating blaze wavelength.

  - declaration:

      - bool	spec_backup (int id);

  - arguments:

      - id : device ID number.

  - return:

        - true: backup success.

        - false: backup fail.

- spec_restore

  - description:

      - reset the device.

  - declaration:

      - bool	 spec_restore (int id);

  - arguments:

      - id : device ID number.

  - return:

        - true: reset success.

        - false: reset fail.

- spec_set_total_steps

  - description:

      - set total steps of  the turret.

  - declaration:

      - bool	 spec_set_total_steps(int id, long steps);

  - arguments:

      - id : device ID number.

      - step : steps to be set.

  - return:

        - true: set steps successfully

        - false: set fail.

- spec_get_total_steps

  - description:

      - get total steps of  the turret.

  - declaration:

      - bool	 spec_get_total_steps(int id, long * steps);

  - arguments:

      - id : device ID number.

      - steps : total steps of  the turret.

  - return:

        - true: get steps successfully.

        - false: get fail.

- spec_set_io_output

  - description:

      - set IO output port number.

  - declaration:

      - bool	 spec_set_io_output (int id, short val);

  - arguments:

      - id : device ID number.

      - val : output port number of  the device.

  - return:

        - true: set success.

        - false:set fail.

- spec_get_io_output

  - description:

      - get IO output port number.

  - declaration:

      - bool	spec_get_io_output (int id, short * val);

  - arguments:

      - id : device ID number.

      - val : output port number of  the device.

  - return:

        - true: get IO number success.

        - false: get fail.

- spec_set_turret

  - description:

      - set the current turret number.

  - declaration:

      - bool	spec_set_turret (int id, int num);

  - arguments:

      - id : device ID number.

      - num : turret number only from 1 to 3.

  - return:

        - true: set success.

        - false: set fail.

- spec_get_turret

  - description:

      - get the current turret number.

  - declaration:

      - bool	spec_get_turret (int id, int * num);

  - arguments:

      - id : device ID number.

      - num : turret number.

  - return:

        - true:  get the number successfully.

        - false:  get the number fail.

- spec_set_zero_pos

  - description:

      - set zero-level position of the grating.The third parameter takes effect only on the spectrometers of the Omni300S model and the HiperS model.

  - declaration:

      - bool	 spec_set_zero_pos (int id, int grating, int mirror，long step);

  - arguments:

      - id : device ID number.

      - grating : the grating number to be set.

      - mirror : 0-front en & front ex, 1-front en & side ex, 2-side en & front ex, 3-side en & side ex.

      - step : motor steps of zero-level position, range from 0 to total steps.

  - return:

        - true: set success.

        - false: set fail.

- spec_get_zero_pos

  - description:

      - get zero-level position of the grating, The third parameter takes effect only on the spectrometers of the Omni300S model and the HiperS model.

  - declaration:

      - bool	spec_get_zero_pos (int id, int grating, int mirror, long* step);

  - arguments:

      - id : device ID number.

      - grating : the grating number to be get.

      - mirror : 0-front en & front ex, 1-front en & side ex, 2-side en & front ex, 3-side en & side ex.

      - step : motor steps of zero-level position.

  - return:

        - true:  get the step successfully.

        - false:  get the step fail.

- spec_set_adjustment

  - description:

      - set device correction coefficients.

  - declaration:

      - bool	spec_set_adjustment(int id, int grating, float val);

  - arguments:

      - id : device ID number.

      - grating : the grating number to be set.

      - val : coefficient parameter.

  - return:

        - true: set success.

        - false: set fail.

- spec_get_adjustment

  - description:

      - get device correction coefficients.

  - declaration:

      - bool	spec_get_adjustment(int id, int grating, float * val);

  - arguments:

      - id : device ID number.

      - grating : the grating number to be get.

      - val : coefficient parameter.

  - return:

        - true:  get coefficient parameter successfully.

        - false:  get coefficient parameter fail.

- spec_get_max_wavelength

  - description:

      - get maximum range of grating.

  - declaration:

      - bool	spec_get_max_wavelength(int id, int grating, float * val);

  - arguments:

      - id : device ID number.

      - grating : the grating number to be get.

      - val : maximum range value.

  - return:

        - true:  get range value successfully.

        - false:  get range value fail.

- spec_adjusting

  - description:

      - Calculate the correction coefficient.

  - declaration:

      - bool	spec_adjusting(int id, int grating, float wave_set, float wave_real, float * result);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - wave_set : Theoretical wavelength value of characteristic peak.

      - wave_real : Measured wavelength value of characteristic peak.

      - result : correction coefficient.

  - return:

a)  true: calculation complete.

b)  false: calculation fail.

- spec_set_grating_info

  - description:

      - set grating groove density and blaze wavelength.

  - declaration:

      - bool	spec_set_grating_info(int id, int grating, int groove, long blaze);

  - arguments:

      - id : device ID number.

      - grating :the grating number.

      - groove : the grating groove density value.

      - blaze : blaze wavelength.

  - return:

        - true: set grating info success.

        - false: set grating info fail.

- spec_get_grating_info

  - description:

      - get grating groove density and blaze wavelength.

  - declaration:

      - bool	spec_get_grating_info(int id, int grating, int * groove, long * blaze);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - groove : the grating groove density value.

      - blaze : blaze wavelength.

  - return:

        - true:  get grating info successfully.

        - false:  get grating info fail.

- spec_set_init_grating

  - description:

      - set the start position of grating during the power-on initialization of device.

  - declaration:

      - bool	spec_set_init_grating(int id, int grating);

  - arguments:

      - id : device ID number.

      - grating : range is 0 to 3, where 0 represents the position at the last shutdown.

  - return:

        - true: set initialization successfully.

        - false: set initialization fail.

- spec_get_init_grating

  - description:

      - get the start position of grating during the power-on initialization of device.

  - declaration:

      - bool	spec_get_init_grating(int id, int * grating);

  - arguments:

      - id : device ID number.

      - grating : range is 0 to 3, where 0 represents the position at the last shutdown..

  - return:

        - true:  get start position successfully.

        - false:  get start position fail.

- spec_set_init_wave

  - description:

      - set the position of the grating when switching the grating.

  - declaration:

      - bool	spec_set_init_wave(int id, int grating, float wave);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - wave : wavelength position of the grating, the unit is nm.

  - return:

        - true: set the position success.

        - false: set the position fail.

- spec_get_init_wave

  - description:

      - get the position of the grating when switching the grating.

  - declaration:

      - bool	spec_get_init_wave(int id, int grating, float * wave);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - wave : wavelength position of the grating, the unit is nm.

  - return:

        - true:  get the position successfully.

        - false:  get the position fail.

- spec_set_dispersion

  - description:

      - set the reverse line dispersion of the device.

  - declaration:

      - bool	spec_set_dispersion(int id, float val);

  - arguments:

      - id : device ID number.

      - val : reverse line dispersion of the device. unit is nm/mm.

  - return:

        - true: set the parameter success.

        - false: set the parameter fail.

- spec_get_dispersion

  - description:

      - get the reverse line dispersion of the device.

  - declaration:

      - bool	spec_get_dispersion(int id, float * val);

  - arguments:

      - id : device ID number.

      - val : reverse line dispersion of the device. unit is nm/mm.

  - return:

        - true:  get the parameter successfully.

        - false:  get the parameter fail.

- spec_set_grating_home

  - description:

      - move the grating to 0 nm position.

  - declaration:

      - bool	spec_set_grating_home(int id);

  - arguments:

      - id : device ID number.

  - return:

        - true: set grating to 0 nm position success.

        - false: set grating to 0 nm position set fail.

- spec_get_grating_count

  - description:

      - Get the total number of grating.

  - declaration:

      - bool	spec_spec_get_grating_count(int id, short * count);

  - arguments:

      - id : device ID number.

      - count : pointer to the total number of grating.

  - return:

        - true: success.

        - false: fail.

- spec_set_grating

  - description:

      - switch the grating.

  - declaration:

      - bool	spec_set_grating(int id, int grating);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

  - return:

        - true: set the grating success(The device will run for a while).

        - false: set the grating fail.

- spec_get_grating

  - description:

      - get the grating number.

  - declaration:

      - bool	spec_get_grating(int id, int * grating);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

  - return:

        - true:  get the grating number successfully.

        - false:  get the grating number fail.

- spec_move_wave

  - description:

      - move the grating to a new position from the current position.

  - declaration:

      - bool	spec_move_wave(int id, float wave);

  - arguments:

      - id : device ID number.

      - wave : wavelength to be moved. The unit is nm.

  - return:

        - true: movement success.

        - false: movement fail because of incorrect ID number or target position is out of range.

- spec_move_to_wave

  - description:

      - move the grating to the wavelength position that is set.

  - declaration:

      - bool	spec_move_to_wave(int id, float wave);

  - arguments:

      - id : device ID number.

      - wave : target wavelength to be set. The unit is nm.

  - return:

        - true: movement success.

        - false: movement fail because of incorrect ID number or target position is out of range.

- spec_get_curr_steps

  - description:

      - get current steps of the grating.

  - declaration:

      - bool	spec_get_curr_steps(int id, float * steps);

  - arguments:

      - id : device ID number.

      - steps : steps of the grating.

  - return:

        - true:  get the steps successfully.

        - false:  get the steps fail.

- spec_move_steps

  - description:

      - move the motor relative to the current position by a specified number of steps.

  - declaration:

      - bool	spec_move_steps(int id, float steps);

  - arguments:

      - id : device ID number.

      - steps : steps to go.

  - return:

        - true:  move steps successfully.

        - false:  move steps fail.

- spec_move_to_steps

  - description:

      - move the motor to a specified position of steps.

  - declaration:

      - bool	spec_move_to_steps (int id, float steps);

  - arguments:

      - id : device ID number.

      - steps : target position of steps.

  - return:

        - true:  move steps successfully.

        - false:  move steps fail.

- spec_get_curr_wave

  - description:

      - get the current wavelength at the position.

  - declaration:

      - bool	spec_spec_get_curr_wave(int id, float * wave);

  - arguments:

      - id : device ID number.

      - wave: the wavelength. The unit is nm.

  - return:

        - true:  get the wavelength successfully.

        - false:  get the wavelength fail.

- spec_set_move_speed

  - description:

      - set the motor speed. Generally, it is not recommended for users to modify it, otherwise it will affect the accuracy of the device's motion results.

  - declaration:

      - bool	spec_set_move_speed(int id, int grating, int speed);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - speed : speed value.

  - return:

        - true: set motor speed successfully.

        - false: set fail.

- spec_get_move_speed

  - description:

      - get the motor speed.

  - declaration:

      - bool	spec_get_move_speed(int id, int grating, int * speed);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - speed : speed value.

  - return:

        - true: get the motor speed successfully.

        - false: get the speed fail.

- spec_set_filter_status

  - description:

      - set the filter wheel status.

  - declaration:

      - bool spec_set_filter_status(int id, int index, bool automatic);

  - arguments:

      - id : device ID number.

      - index:0-Outside, 1-Inside

      - automatic :

        - true: The filter wheel rotates automatically according to the wavelength.

        - false: The filter wheel rotates manually.

  - return:

      - true:success

      - false:failed.

- spec_get_filter_status

  - description:

      - get the filter wheel status.

  - declaration:

      - void	spec_get_filter_status(int id, int index, bool * automatic);

  - arguments:

      - id : device ID number.

      - index:0-Outside, 1-Inside

      - automatic :

        - true: The filter wheel rotates automatically according to the wavelength.

        - false: The filter wheel rotates manually.

  - return:

      - void.

- spec_set_filter

  - description:

      - select the position of the filter wheel manually.

  - declaration:

      - bool	spec_set_filter(int id, int index, int which);

  - arguments:

      - id: device ID number.

      - index:0-Outside, 1-Inside

      - which: position index number of filter wheel from 1 to 6.

  - return:

        - true: set the position of filter wheel successfully.

        - false: set the position fail.

- spec_get_filter

  - description:

      - get the position of the filter wheel.

  - declaration:

      - bool	spec_get_filter(int id, int index, int * which);

  - arguments:

      - id : device ID number.

      - index:0-Outside, 1-Inside

      - which: position index number of filter wheel from 1 to 6.

  - return:

        - true: get the position of filter wheel successfully.

        - false: get the position fail.

- spec_set_filter_home

  - description:

      - set the filter wheel to zero-position.

  - declaration:

      - bool	spec_set_filter_home(int id, int index);

  - arguments:

      - id : device ID number.

      - index:0-Outside, 1-Inside

  - return:

        - true: set the zero-position successfully.

        - false: set fail.

- spec_set_filter_model

  - description:

      - set the type of filter wheel model.

  - declaration:

      - bool	spec_set_filter_model(int id, const char * model);

  - arguments:

      - id : device ID number.

      - model : filter wheel model, SD or SR-SD.

  - return:

        - true: set filter model successfully.

        - false: set model fail.

- spec_get_filter_model

  - description:

      - get the type of filter wheel model.

  - declaration:

      - bool	spec_get_filter_model(int id, char * model);

  - arguments:

      - id : device ID number.

      - model : filter wheel model, SD or SR-SD.

  - return:

        - true: get filter model successfully.

        - false: get model fail.

- spec_set_filter_limit

  - description:

      - set the working wavelength of each position of thefilter wheel.

  - declaration:

      - bool	spec_set_filter_limit(int id, int grating, short index, double val);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - index : position index number of filter wheel from 1 to 6.

      - val : working wavelength. The unit is nm.

  - return:

a)  true: set the wavelength successfully.

b)  false: set fail because of device ID error or position index number error or grating number error.

- spec_get_filter_limit

  - description:

      - get the working wavelength of each position of thefilter wheel.

  - declaration:

      - bool	spec_get_filter_limit(int id, int grating, short index, double * val);

  - arguments:

      - id : device ID number.

      - grating : the grating number.

      - index : position index number of filter wheel from 1 to 6.

      - val : working wavelength. The unit is nm.

  - return:

        - true: get the working wavelength successfully.

        - false: get fail because of device ID error or position index number error or grating number error.

- spec_set_exit_port

  - description:

      - set exit port of device.

  - declaration:

      - bool	spec_set_exit_port(int id, bool front);

  - arguments:

      - id : device ID number.

      - front : exit port identifier

        - true: direct exit.

        - false: side exit.

  - return:

        - true: set exit port successfully.

        - false: set exit port fail.

- spec_get_exit_port

  - description:

      - get exit port of device.

  - declaration:

      - bool	spec_get_exit_port(int id, bool * front);

  - arguments:

      - id : device ID number.

      - front : exit port identifier

        - true: direct exit.

        - false: side exit.

  - return:

        - true: get exit port successfully.

        - false: get exit port fail.

- spec_set_side_exit_pos

  - description:

      - set the steps of the side exit position.

  - declaration:

      - bool	spec_set_side_exit_pos(int id, int step);

  - arguments:

      - id : device ID number.

      - step : steps value to be set from -3200 to 0.

  - return:

        - true: set the steps successfully.

        - false: set the steps fail.

- spec_get_side_exit_pos

  - description:

      - get the steps of the side exit position.

  - declaration:

      - bool	spec_get_side_exit_pos(int id, int * step);

  - arguments:

      - id : device ID number.

      - step : steps value from -3200 to 0.

  - return:

        - true: get the steps successfully.

        - false: get the steps fail.

- spec_set_entrance_port

  - description:

      - set entrance port of device.

  - declaration:

      - bool	spec_set_entrance_port (int id, bool front);

  - arguments:

      - id : device ID number.

      - front : entrance port identifier

        - true: direct entrance.

        - false: side entrance.

  - return:

        - true: set entrance port successfully.

        - false: set entrance port fail.

- spec_get_entrance_port

  - description:

      - get entrance port of device.

  - declaration:

      - bool	spec_get_entrance_port(int id, bool * front);

  - arguments:

      - id : device ID number.

      - front : entrance port identifier

        - true: direct entrance.

        - false: side entrance.

  - return:

        - true: get entrance port successfully.

        - false: get entrance port fail.

- spec_set_side_entrance_pos

  - description:

      - set the steps of the side entrance position.

  - declaration:

      - bool	spec_set_side_entrance_pos(int id, int step);

  - arguments:

      - id : device ID number.

      - step : steps value to be set from -3200 to 0.

  - return:

        - true: set the steps successfully.

        - false: set the steps fail.

- spec_get_side_entrance_pos

  - description:

      - get the steps of the side entrance position.

  - declaration:

      - bool	spec_get_side_entrance_pos(int id, int * step);

  - arguments:

      - id : device ID number.

      - step : steps value from -3200 to 0.

  - return:

        - true: get the steps successfully.

        - false: get the steps fail.

- spec_set_slit_factor

- spec_get_slit_factor

- spec_set_slit_width

  - description:

      - set the automatic slit width.

  - declaration:

      - bool	spec_set_slit_width(int id, int index, int um);

  - arguments:

      - id : device ID number.

      - index : the automatic slit number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - um : slit width with the unit of um.The value must be a multiple of 10um and the minimum value is 10um.

  - return:

        - true: set the automatic slit width successfully.

        - false: set the automatic slit width fail that may be caused by device ID error or wrong automatic slit number or invalid slit width value.

- spec_get_slit_width

  - description:

      - get the automatic slit width.

  - declaration:

      - bool	spec_get_slit_width(int id, int index, int * um);

  - arguments:

      - id : device ID number.

      - index : the automatic slit number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - um : slit width with the unit of um.

  - return:

        - true: get the automatic slit width successfully.

        - false: get the automatic slit width fail.

- spec_set_slit_bandpass

  - description:

      - set the automatic slit bandwidth.

  - declaration:

      - bool	spec_set_slit_bandpass(int id, int index, float nm);

  - arguments:

      - id : device ID number.

      - index : the grating number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - nm : slit bandwidth value to be set. The unit is nm.

  - return:

        - true: set the automatic slit bandwidth successfully.

        - false: set the automatic slit width fail that may be caused by device ID error or wrong automatic slit number or invalid slit bandwidth value.

- spec_get_slit_bandpass

  - description:

      - get the automatic slit bandwidth.

  - declaration:

      - bool	spec_get_slit_bandpass(int id, int index, float * nm);

  - arguments:

      - id : device ID number.

      - index: the automatic slit number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - nm : slit bandwidth value. The unit is nm.

  - return:

        - true: get the automatic slit bandwidth successfully.

        - false: get the automatic slit bandwidth fail.

- spec_set_slit_zero_pos

  - description:

      - set the automatic slit zero position.

  - declaration:

      - bool 	spec_set_slit_zero_pos(int id, int index, int um);

  - arguments:

      - id : device ID number.

      - index : the grating number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - nm : slit width to zero position. The unit is um and range is from -100 to 100um.

  - return:

        - true: set the automatic slit zero position successfully.

        - false: set fail that may be caused by device ID error or wrong slit number or invalid slit width value.

- spec_get_slit_zero_pos

  - description:

      - get the automatic slit zero position.

  - declaration:

      - bool	spec_get_slit_zero_pos(int id, int index, int * um);

  - arguments:

      - id : device ID number.

      - index : the automatic slit number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - um : slit width to zero position. The unit is um and range is from -100 to 100um.

  - return:

        - true: get the automatic slit zero position successfully.

        - false: get zero position fail.

- spec_set_slit_home

  - description:

      - automatic slit reset.

  - declaration:

      - bool	spec_set_slit_home(int id, int index);

  - arguments:

      - id : device ID number.

      - index : automatic slit position number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

  - return:

        - true: automatic slit reset success.

        - false: reset fail.

- spec_set_slit_model

  - description:

      - set the automatic slit model.

  - declaration:

      - bool	spec_set_slit_model(int id, int index, char * model);

  - arguments:

      - id : device ID number.

      - index : automatic slit position number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - model : automatic slit model, SLIT or SLIT-I24 or SR-SLIT.

  - return:

        - true: set slit model successfully.

        - false: set slit model fail.

- spec_get_slit_model

  - description:

      - get the automatic slit model.

  - declaration:

      - bool	spec_get_slit_model(int id, int index, char * model);

  - arguments:

      - id : device ID number.

      - index : automatic slit position number.

0:side entrance       1:side exit

2:direct exit          3:direct entrance

      - model : automatic slit model, SLIT or SLIT-I24 or SR-SLIT.

  - return:

        - true: get the automatic slit model successfully.

        - false: get the model fail.

- spec_set_motor_steps

  - description:

      - set the absolute steps of the stepper motor.

  - declaration:

      - bool	spec_set_motor_steps(int id, int index, long steps);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor, 3 to 13.

      - steps :

  - return:

        - true: set motor steps successfully.

        - false: set steps fail.

- spec_get_motor_steps

  - description:

      - get the absolute steps of the stepper motor.

  - declaration:

      - bool	spec_get_motor_steps(int id, int index, long * steps);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor.

      - steps :

  - return:

        - true: get motor steps successfully.

        - false: get steps fail.

- spec_set_motor_home

  - description:

      - set the stepper motor to the start-point position.

  - declaration:

      - bool	spec_set_motor_home(int id, int index);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor, 2 to 13.

  - return:

        - true: set the motor to the start-point successfully.

        - false: set fail.

- spec_set_motor_speed

  - description:

      - set the stepper motor speed.

  - declaration:

      - bool	spec_set_motor_speed(int id, int index, int speed);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor, 3 to 13.

      - speed : speed value of motor from 10 to 2000. This parameter is not saved in the device.

  - return:

        - true: set the stepper motor speed successfully.

        - false: set fail.

- spec_get_motor_speed

  - description:

      - get the stepper motor speed.

  - declaration:

      - bool	spec_get_motor_speed(int id, int index, int * speed);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor.

      - speed : speed value of motor from 10 to 2000. This parameter is not saved in the device.

  - return:

        - true: get the stepper motor speed successfully.

        - false: get speed fail.

- spec_set_motor_home_dir

  - description:

      - set the moving direction of the stepper motor.

  - declaration:

      - bool	spec_set_motor_home_dir(int id, int index, int dir);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor.

      - dir : direction of the stepper motor

        - 1: move in the positive direction.

        - 0: move in the opposite direction.

  - return:

        - true: set the moving direction successfully.

        - false: set fail.

- spec_get_motor_home_dir

  - description:

      - get the moving direction of the stepper motor.

  - declaration:

      - bool	spec_get_motor_home_dir(int id, int index, int * dir);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor.

      - dir : direction of the stepper motor

        - 1: move in the positive direction.

        - 0: move in the opposite direction.

  - return:

        - true: get the moving direction successfully.

        - false: get fail.

- spec_set_motor_total_steps

  - description:

      - set the total number of steps of the stepper motor.

  - declaration:

      - bool	spec_set_motor_total_steps(int id, int index, long total);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor from 2 to 13.

      - total : total steps from 0 to 2147483647.

  - return:

        - true: set the total number of steps successfully.

        - false:set fail.

- spec_get_motor_total_steps

  - description:

      - get the total number of steps of the stepper motor.

  - declaration:

      - bool	spec_get_motor_total_steps(int id, int index, long * total);

  - arguments:

      - id : device ID number.

      - index : The serial number of the stepper motor from 2 to 13.

      - total : total steps.

  - return:

        - true: get the total number of steps successfully.

        - false: get fail.

- spec_set_shutter_status

  - description:

      - set the opening and closing state of the shutter.

  - declaration:

      - bool	spec_set_shutter_status(int id, int port, bool status);

  - arguments:

      - id : device ID number.

      - port : Shutter installation position

- 0: side in. 		   1: direct in.

- 2:side out			2:direct out

      - status : state of the shutter

        - true: open.			false: close.

  - return:

        - true: set shutter status successfully.

        - false: set fail.

- spec_get_shutter_status

  - description:

      - get the opening and closing state of the shutter.

  - declaration:

      - bool	spec_get_shutter_status(int id, int port, bool * status);

  - arguments:

      - id : device ID number.

      - port : Shutter installation position.

- 0: side in. 		   1: direct in.

- 2:side out			2:direct out

      - status : state of the shutter

        - true: open.		false: close.

  - return:

        - true: get shutter status successfully.

        - false: get fail.

- spec_is_setup_filter

  - description:

      - check if the filter wheel is installed.

  - declaration:

      - bool	spec_is_setup_filter(int id, int index, bool * setup);

  - arguments:

      - id : device ID number.

      - index:0-Outside, 1-Inside

      - setup : installation status

        - true: installed.			false: not installed.

  - return:

        - true: check success.

        - false: check fail.

- spec_setup_filter

  - description:

      - set whether the device needs to install a filter wheel, only for side port installation.

  - declaration:

      - bool	spec_setup_filter(int id, int index, bool setup);

  - arguments:

      - id : device ID number.

      - index:0-Outside, 1-Inside

- spec_is_setup_mirror

  - description:

      - check if the mirror is installed.

  - declaration:

      - bool	spec_is_setup_mirror(int id, int index, bool * setup);

  - arguments:

      - id : device ID number.

      - index : Mirror installation position

        - 0: exit.   		    1: entrance.

      - setup :

        - true: installed.			false: not installed.

  - return:

        - true: check success.

        - false: check fail.

- spec_setup_mirror

  - description:

      - set whether the divice needs to install a mirror.

  - declaration:

      - bool	spec_setup_mirror(int id, int index, bool setup);

  - arguments:

      - id : device ID number.

      - index : Mirror installation position.

        - 0: Exit.		    1: entrance.

      - setup :

        - true: need to install.		  false: no need to install.

  - return:

        - true: set success.

        - false: set fail.

- spec_is_setup_slit

  - description:

      - check if the slit is installed.

  - declaration:

      - bool	spec_is_setup_slit(int id, int index, bool * setup);

  - arguments:

      - id : device ID number.

      - index : slit installation position

        - 0: side entrance.			    1: side exit.

        - 2: direct exit.			        3: direct entrance.

      - setup :

        - true: installed.			   false: not installed.

  - return:

        - true: check success.

        - false: check fail.

- spec_setup_slit

  - description:

      - set whether the device needs to install a slit.

  - declaration:

      - bool	spec_setup_slit(int id, int index, bool setup);

  - arguments:

      - id : device ID number.

      - index : Slit installation position.

        - 0: side entrance.		    1: side exit.

        - 2: direct exit.			    3: direct entrance.

      - setup :

        - true: installed.		false: not installed.

  - return:

        - true: set success.

        - false: set fail.

- spec_is_setup_shutter

  - description:

      - check if the shutter is installed.

  - declaration:

      - bool	spec_is_setup_shutter(int id, int index, bool * setup);

  - arguments:

      - id : device ID number.

      - index : shutter installation position.

        - 0: side entrance.			    1: direct entrance.

        - 2: side exit.			    	 3: direct exit.

      - setup :

        - true: installed.			false: not installed.

  - return:

        - true: check success.

        - false: check fail.

- spec_setup_shutter

  - description:

      - set whether the device needs to install a shutter.

  - declaration:

      - bool	spec_setup_shutter(int id, int index, bool setup);

  - arguments:

      - id : device ID number.

      - index : Shutter installation position.

        - 0: side entrance.			    1: direct entrance.

        - 2: side exit.			    	 3: direct exit.

      - setup :

        - true: installed.  	false: not installed.

  - return:

        - true: set success.

        - false: set fail.

- spec_set_correct_params

  - description:

      - set the grating correction coefficients. Different device versions have different parameter types. Please refer to demo project for specific usage methods.

  - declaration:

      - bool	spec_set_correct_params(int id, int turret, int grating, int code, void * param);

  - arguments:

      - id : device ID number.

      - turret : turret number.(1~3)

      - grating : the grating number.(1~3)

      - code : correction coefficient identifier 0~16 and its meaning.

- Code							V3.4and later | afterV3.4

- 0: focal length								float	float

- 1: inclusion angle							float	float

- 2: CCD tilt angle							float	float

- 3: grating factor 1  						double	float

- 4: grating factor 2  						double	float

- 5: grating factor 3							double	float

- 6: grating factor 4							double	float

- 7: grating factor 5							double	float

- 8: cosine factor—amplitude				double	float

- 9: cosine factor—phase					double	float

- 10: cosine factor—frequency				double	float

- 11: start position							int		int

- 12: start position offset					short	int

- 13: grating groove density				double	float

- 14: grating blaze wavelength				double	-

- 15: front entrance & direct exit offset		short	short

- 16: front entrance & side exit offset		short	short

- 17: side entrance & direct exit offset		short  short

- 18: side entrance & & side exit offset		short	short

      - param : factor value to be set.

  - return:

        - true: set factor value successfully.

        - false:set fail.

- spec_get_correct_params

  - description:

      - get the grating correction coefficients. Different device versions have different parameter types. Please refer to demo project for specific usage methods.

  - declaration:

      - bool	spec_get_correct_params(int id, int Turret, int grating, int code, void * param);

  - arguments:

      - id : device ID number.

      - turret: the turret number.(1~3)

      - grating : the grating number.(1~3)

      - code : correction coefficient identifier 0~16 and its meaning.

- Code							V3.4and later | afterV3.4

- 0: focal length								float	float

- 1: inclusion angle							float	float

- 2: CCD tilt angle							float	float

- 3: grating factor 1  						double	float

- 4: grating factor 2  						double	float

- 5: grating factor 3							double	float

- 6: grating factor 4							double	float

- 7: grating factor 5							double	float

- 8: cosine factor—amplitude				double	float

- 9: cosine factor—phase					double	float

- 10: cosine factor—frequency				double	float

- 11: start position							int		int

- 12: start position offset					short	int

- 13: grating groove density				double	float

- 14: grating blaze wavelength				double	-

- 15: front entrance & direct exit offset		short	short

- 16: front entrance & side exit offset		short	short

- 17: side entrance & direct exit offset		short  -

- 18: side entrance & & side exit offset		short	-

      - param : factor value to be set.

  - return:

        - true: get factor value successfully.

        - false: get fail.

- spec_wave_to_step

  - description:

      - convert the wavelength to the motor steps.

  - declaration:

      - float	spec_wave_to_step(int id, float wave);

  - arguments:

      - id : device ID number.

      - wave : wavelength to move. the unit is nm.

  - return:

        - 0: Calculation failed.

        - Non-zero value: The position where the motor will move.

- spec_pixels_to_waves

  - description:

      - convert pixel position to wavelength in batches.

  - declaration:

      - bool	spec_pixels_to_waves(int id, int turret, int grating, float center_wave, int width, int count, int bin, float * result);

  - arguments:

      - id : device ID number.

      - turret: Turret number, 1 - 3.

      - grating: Grating number, 1 - spec_get_grating_count

      - center_wave : the unit is nm.

      - width : pixel width. The unit is um.

      - count : CCD total pixels.

      - bin : Multiple of horizontal binning, default to1.

      - result : wavelength array with the same size as the total number of CCD pixels.

  - return:

true: Success.

false: Fail,none parameters.

- spec_set_init_peripherals

description:Firmware version at least 3.4

Set whether the peripheral is initialized after power-on

declaration:

bool	spec_set_init_peripherals(int id, unsigned short peripherals);

arguments:

id : device ID number.

peripherals: peripheral code, Can be used in combination.

0x0001 Filter

0x0002 Export mirror

0x0004 Slit - Side in

0x0008 Slit - Side out

0x0010 Slit - Front out

0x0020 Slit - Front in

0x0040 Entrance mirror

0x0080 Filter Inside

0x0100 Shutter 1

0x0200 Shutter 2

0x0400 Focus mirror (HiperS Model only)

0x0800 Reserve (HiperS Model only)

0x1000 Diaphragm at the side entrance (HiperS Model only)

0x2000 Diaphragm at the front entrance (HiperS Model only)

return:

true: Success.

false: Fail.

- spec_get_init_peripherals

description:Firmware version at least 3.4

Read peripherals that need to be initialized at power-up.

declaration:

bool	spec_get_init_peripherals(int id, unsigned short * peripherals);

arguments:

id : device ID number.

peripherals: peripheral code, Can be used in combination.

0x0001 Filter

0x0002 Export mirror

0x0004 Slit - Side in

0x0008 Slit - Side out

0x0010 Slit - Front out

0x0020 Slit - Front in

0x0040 Entrance mirror

0x0080 Filter Inside(Motor No. 9)

0x0100 Shutter 1

0x0200 Shutter 2

0x0400 Focus mirror (HiperS Model only)

0x0800 Reserve (HiperS Model only)

0x1000 Diaphragm at the side entrance (HiperS Model only)

0x2000 Diaphragm at the front entrance (HiperS Model only)

return:

true: Success.

false: Fail.

- spec_set_peripherals_init_pos

description: Firmware version at least 3.4

Set peripheral power-on initialization position.

declaration:

bool	spec_set_peripherals_init_pos(int id, unsigned short peripheral, unsigned char pos);

arguments:

id : device ID number.

peripherals: peripheral code, Can’t be used in combination.

0x0001 Filter

0x0002 Export mirror

0x0004 Slit - Side in

0x0008 Slit - Side out

0x0010 Slit - Front out

0x0020 Slit - Front in

0x0040 Entrance mirror

0x0080 Motor No. 9

0x0100 Shutter 1

0x0200 Shutter 2

0x0400 Focus mirror (HiperS Model only)

0x0800 Reserve (HiperS Model only)

0x1000 Diaphragm at the side entrance (HiperS Model only)

0x2000 Diaphragm at the front entrance (HiperS Model only)

Pos: 0 - Zero position, 1 - Position when last shut down.

return:

true: Success.

false: Fail.

- spec_get_peripherals_init_pos

description: Firmware version at least 3.4

Read the initial position of the peripheral at power-up.

declaration:

bool	spec_get_peripherals_init_pos(int id, unsigned short peripheral, unsigned char * pos);

arguments:

id : device ID number.

peripherals: peripheral code, Can’t be used in combination.

0x0001 Filter

0x0002 Export mirror

0x0004 Slit - Side in

0x0008 Slit - Side out

0x0010 Slit - Front out

0x0020 Slit - Front in

0x0040 Entrance mirror

0x0080 Motor No. 9

0x0100 Shutter 1

0x0200 Shutter 2

0x0400 Focus mirror (HiperS Model only)

0x0800 Reserve

0x1000 Diaphragm at the side entrance (HiperS Model only)

0x2000 Diaphragm at the front entrance (HiperS Model only)

pos: 0 - Zero position, 1 - Position when last shut down.

return:

true: Success.

false: Fail.

- spec_set_trig_out_interval

description:Firmware version at least 3.4

After setting the moving range and interval steps, the monochromator will send out a pulse every interval step during the absolute movement (wavelength) process without feedback.

declaration:

bool	spec_set_trig_out_interval(int id, unsigned short steps);

arguments:

id :  device ID number.

steps : interval steps, in 0-65535.

return:

true: Success.

false: Fail.

- spec_get_trig_out_interval

description:Firmware version at least 3.4

Read the number of interval steps at which the monochromator sends out pulses during the movement.

declaration:

bool	spec_get_trig_out_interval(int id, unsigned short * steps);

arguments:

id :  device ID number.

steps : interval steps.

return:

true: Success.

- false: Fail.

- spec_set_trig_mode

description:Firmware version at least 3.4

Start or stop trigger-in mode.

declaration:

bool	spec_set_trig_mode(int id, bool start);

arguments:

id :  device ID number.

start : trigger-in mode:

true:start.

false:stop.

return:

true: Success.

- false: Fail.

- spec_set_trig_in_interval

description:Firmware version at least 3.4

Set the interval wavelength moved by each pulse signal.

declaration:

bool	spec_set_trig_in_interval(int id, float nm);

arguments:

id :  device ID number.

nm: interval wavelength, unit - nm

return:

true: Success.

- false: Fail.

- spec_get_trig_in_interval

description:Firmware version at least 3.4

Read the interval wavelength moved by each pulse signal.

declaration:

bool	spec_get_trig_in_interval(int id, float * nm);

arguments:

id :  device ID number.

nm: interval wavelength, unit - nm.

return:

true: Success.

- false: Fail.

- spec_range_move

description:Firmware version at least 3.4

After setting the moving range and moving interval, the monochromator will move within the range.

declaration:

bool	spec_range_move(int id, float from_nm, float to_nm, float step_nm,int nm_s);

arguments:

id :  device ID number.

from_nm: Start of moving wavelength, unit - nm.

to_nm: interval wavelength, unit - nm.

step_nm: End of moving wavelength, unit - nm.

nm_s: movement speed, unit:nm/s, less than the value read from spec_get_move_speed.

return:

true: Success.

- false: Fail

- spec_range_move2

description:Firmware version at least 3.4

Set the monochromator range move, wait for the delay (ms) after each interval, and then send a pulse signal outwards.

declaration:

bool	spec_range_move2(int id, float from_nm, float to_nm, float step_nm,int delay_ms);

arguments:

id :  device ID number.

from_nm: Start of moving wavelength, unit - nm.

to_nm: interval wavelength, unit - nm.

step_nm: End of moving wavelength, unit - nm.

delay_ms: interval

return:

true: Success.

- false: Fail

- spec_set_ccd_mode

description:

Set the spectrometer to be used with CCD.

declaration:

bool	spec_set_ccd_mode(int id, bool ccd_mode);

arguments:

id :  device ID number.

ccd_mode: To be used with CCD or not.

return:

true: Success.

- false: Fail

- Note: In firmware version 3.3 and earlier, the central wavelength of the current grating position in CCD mode is recorded in the local memory. After the software is restarted, the central wavelength of the CCD at the current grating position cannot be read; In firmware version 3.4 and above, the CCD central wavelength of the current grating position is recorded in the memory of the spectrometer, and the CCD central wavelength of the current grating position can be read in real time.

- spec_get_ccd_mode

description:

Read whether the work mode of spectrometer is used with CCD.

declaration:

bool	spec_get_ccd_mode(int id);

arguments:

id :  device ID number.

return:

true: Work with ccd.

- false: Normal mode.

- spec_init_spectral_splice

description:

Initialize the splicing spectrum parameters and calculate the center wavelength required for splicing spectrum.

declaration:

bool	spec_init_spectral_splice(int id, float pixel_width, int pixel_count, int bin_x, float from, float to, float ref,int edge, float op, float* center_waves, int* waves_count);

arguments:

[IN] id : device ID number.

[IN] pixel_width: Width of pixels.

[IN] pixel_count: Number of pixels.

[IN] bin_x: Horizontal direction binning multiple, default 1.

[IN] from: begin of wavelength.

[IN] to: end of wavelength.

[IN] ref: reference wavelength, default 500 nm.

[IN] edge: The number of edge pixels, default 50.

[IN] op: overlap percent, range is 0.1 - 0.5.

[OUT] center_waves: Center wavelength buffer, initialized before use, size of to - from is enough

[OUT] waves_count: Actual Legnth of center wavelength buffer.

return:

true: Success.

- false: Fail

- spec_spectral_splice

description:

Start Splicing wavelength and spectral data, please initialize parameters first.

declaration:

bool	spec_spectral_splice(int id, float* x1, float* y1, int size1,float* x2, float* y2, int size2, float* x3, float* y3, int* size3);

arguments:

[IN] id : device ID number

[IN] x1: Wavelength data before splicing

[IN] y1: Spectral data befor splicing

[IN] size1: Legnth of Wavelength data before splicing

[IN] x2: Wavelength data to be spliced

[IN] y2: Spectral data to be spliced

[IN] size2: Legnth of Wavelength data to be spliced

[OUT] x3: Wavelength data after splicing,need to be initialized before use,size of size1+size2 is enough

[OUT] y3: Spectral data after splicing,need to be initialized before use,size of size1+size2 is enough

[OUT] size3: Actual Legnth of Wavelength data after splicing

return:

true: Success.

- false: Fail

- spec_set_user_data

description:Firmware version at least 3.4

Write user data into EEPROM, The effective buffer size is 8190 kb.

declaration:

bool	spec_set_user_data(int id, short offset, const unsigned char * data, short len);

arguments:

id :  device ID number.

offset: memory offset.

data: user data.

len: data length.

return:

true: Success.

- false: Fail

- spec_get_user_data

description:Firmware version at least 3.4

Read the user data from EEPROM.

declaration:

bool	spec_get_user_data(int id, short offset, unsigned char * data, short len);

arguments:

id :  device ID number.

offset: memory offset.

data: user data.

len: data length.

return:

true: Success.

false: Fail

- spec_set_diaphragm

description:

Switch the diaphragm position, Only HiperS model spectrograph is available.

declaration:

bool	spec_set_diaphragm(int id, int which, int where);

arguments:

id : device ID number.

which: which entrance,0-side,1-front.

where: the position of diaphragm,range is: 0~10, 0 for reset.

return:

true: Success.

false: Fail

- spec_get_diaphragm

description:

Read the diaphragm position, Only HiperS model spectrograph is available.

declaration:

bool	spec_get_diaphragm(int id, int which, int* where);

arguments:

id : device ID number.

which: which entrance,0-side,1-front.

where: the position of diaphragm.

return:

true: Success.

false: Fail

- spec_set_diaphragm_steps

description:

Set the motor steps at the target position of the diaphragm, Only HiperS model spectrograph is available.

declaration:

bool	spec_set_diaphragm_steps(int id, int which, int where, long steps);

arguments:

id : device ID number.

which: which entrance,0-side,1-front.

where: the position of diaphragm, range is: 1~10, 0 for reset, can not be modified.

steps: motor steps.

return:

true: Success.

false: Fail

- spec_get_diaphragm_steps

description:

Get the motor steps at the target position of the diaphragm, Only HiperS model spectrograph is available.

declaration:

bool	spec_get_diaphragm_steps(int id, int which, int where, long* steps);

arguments:

id : device ID number.

which: which entrance,0-side,1-front.

where: the position of diaphragm, range is: 1~4.

steps: motor steps.

return:

true: Success.

false: Fail

- spec_set_focus_mirror

description:

Switch the focus mirror position, Only HiperS model spectrograph is available.

declaration:

bool	spec_set_focus_mirror(int id, int where);

arguments:

id :  device ID number.

where: focus position, range is: 0~12.

return:

true: Success.

false: Fail

- spec_get_focus_mirror

description:

Read the focus mirror position, Only HiperS model spectrograph is available.

declaration:

bool	spec_get_focus_mirror(int id, int* where);

arguments:

id :  device ID number.

where: focus position.

return:

true: Success.

false: Fail

- spec_set_focus_mirror_steps

description:

Set the motor steps at the target position of the focus mirror, Only HiperS model spectrograph is available.

declaration:

bool	spec_set_focus_mirror_steps(int id, int where, long steps);

arguments:

id :  device ID number.

where: the position of the focus mirror, range is: 1~12, 0 for reset, can not be modified.

steps: motor steps.

return:

true: Success.

false: Fail

- spec_get_focus_mirror_steps

description:

Get the motor steps at the target position of the focus mirror, Only HiperS model spectrograph is available.

declaration:

bool	spec_get_focus_mirror_steps(int id, int where, long* steps);

arguments:

id :  device ID number.

where: the position of the focus mirror, range is: 1~12.

steps: motor steps.

return:

true: Success.

false: Fail

Motor Information

- #2: Filter-Outside(Normal)

- #3:Export

- #4:Slit1(Side-In)

- #5:Slit2(Side-Out)

- #6:Slit3(Front-Out)

- #7:Slit4(Front-In)

- #8:Entrance

- #9:Filter-Inside (Omni320S Model only)

- #10:Focus mirror (HiperS Model only)

- #12:Diaphragm-Side (HiperS Model only)

- #13:Diaphragm-Front (HiperS Model only)

- Error information

E00: It indicates that the communication between the device and the program has not been established.

E01: Communication error, illegal control command or communication timeout.

E02: EEPROM error, usually because the instrument has not set the parameter. Send parameter setting commands.

E03:The grating parameters are set incorrectly or not set.

E04: The parameter of the filter wheel is entered incorrectly.

E05: The filter wheel command is enabled, but the device is not actually installed or the device is malfunctioning.

E06: Except for the grating parameters, the parameter setting is out of range.

E07: For stepping motors 3~9, before performing positioning action, get position command or movement command are sent ; or performing movement to zero-position command, but the location switch is faulty.

E08: Speed setting exceeds upper limit.

E09: Command to switch the device to direct entrance/exit port or side entrance/exit port, but there is actually no equipment installed or equipment failure.

E10: Rough location switch of grating turret failure.

E11: Accurate location switch of grating turret failure.

E12: The automatic slit command is enabled, but the device is not actually installed or the device is malfunctioning.

E13:Unexpected disconnection(Please check the power or usb cable. If any function is called again after this error occurs, an E00 error message will be replied).

Note: When the connection is complete, if the function call returns a failure, and_ Get it_ ErrorIf the error information read contains "E13" or "E00", it indicates that the hardware connection is abnormal;

When USB is lost, the DLL will automatically reconnect to the spectrometer USB. If the DLL fails to reconnect to the USB for 1 minute, the connection will be stopped. At this time, there may be a problem with the USB or the spectrometer.

It takes about 3-5 seconds for the hardware to automatically connect to the SDK from being identified by the device manager after reconnection.